Monitoring of Thermally Initiated Reactions On-line by Electrospray Ionization Mass Spectrometry

Jens Griep-Raming and Jürgen O. Metzger

University of Oldenburg, Department of Chemistry, P.O. Box 2503, 26111 Oldenburg, Germany, griep@chemie.uni-oldenburg.de

Overview

- A method to investigate thermally initiated organic reactions by electrospray ionization (ESI) mass spectrometry (MS) is described.

- For this purpose a new type of ESI source, which allows to heat the ESI spray capillary directly, has been developed.

- Two examples are given which demonstrate the general applicability of this method for the investigation of solution phase organic reactions.

Introduction
The determination of reaction mechanisms and kinetics is an important task in organic chemistry. However, the methods to determine especially transient intermediates in reaction solutions are limited. In order to develop a new method to detect reactive intermediates in solution, we have designed a new type of ESI source to investigate thermally initiated reactions on-line. Our source design allows to heat the spray capillary directly, thus enabling us to acquire mass spectra of heated reaction solutions.

Methods
The basic design of our ESI source is shown in figure 1. The source consists of a fused silica spray capillary  in which a platinum wire  was inserted. This assembly is surrounded by standard s.s. HPLC tubing, covered by heating tape.

Figure 1 Detail drawing of the ESI source with heatable spray capillary

The heating assembly is attached to the sample inlet, sheath gas inlet and ESI voltage by the use of standard HPLC consumable, such as PEEK tees, crosses, fittings and sleeves. Figure 2 shows the complete apparatus. Solutions flowing through the ESI capillary can be heated by applying a voltage to the heating tape. The resulting temperatures are measured using the incorporated thermocouple.

Figure 2 Assembly of the ESI source

Results
Using this ESI-source, we have investigated the thermal dissociation of several compounds. Two examples are given here.

2,2'-Azobis-(2-(2-imidazolinyl)propane

Azo compound  1  is decomposed thermally to give two 2-(2-imidazolinyl)propyl radicals (2) and nitrogen (scheme 1). The radical  2  can react in different ways, i.e. radical combination, disproportionation or addition to other compounds (e.g. oxygen)

Scheme 1 Thermal dissociation of azo compound 1

Unfortunately the signal of radical 2 at m/z 112 is interfered by the first isotope peak of 3. This makes the accurate measurement of the concentration of 2 difficult. By MS/MS experiments after CID radical 2 and compound 3 can be distinguished unambiguously.  Radical 2 can dissociate by radical induced loss of methyl, while compound 3 cannot. Thus 2 can be indetified by the ion transition m/z 112 -> m/z 97 (spectrum shown in figure 3). Figure 4 shows the measured temperature dependence of the radical 2, observed by the mentioned CID experiment. The measurements are in good agreement with the calculated radical concentration.

Figure 3 MS/MS mass spectrum after CID of m/z 112 ([M+H]+ of radical 2)

Figure 4 Temperature dependence of the signal of radical 2 (observed by the transition m/z 112 -> m/z 97 after CID) (solvent: water, initial concentration of 1 110-4 moll-1)

Trityl Dimer

The trityl dimer dissociates to give two trityl radicals (scheme 1). With the use of dichloromethane/ acetonitrile (50:50, v/v with addition of 1 mmol/l lithium trifluoromethanesulfonate) as the solvent it is possible to oxidize radicals to give cations1,  2.

Scheme 2 Thermal dissociation of the trityl dimer

Spectra acquired under these conditions exhibit only one peak at m/z 243 which is contributed to the trityl cation, generated by electrolytic oxidation of the radical. The trityl dimer cannot be oxidized by this method. The temperature dependence of the trityl signal has been recorded 
Figure 5 shows the measured temperature dependence correlated to the calculated trityl concentration in solution.

Figure 5 	Temperature dependence of the trityl radical signal at m/z 243 compared to the calculated trityl concentration in solution (solvent: dichloromethane/acetonitrile 50:50 with addition of 1mmol/l lithium trifluoromethanesulfonate, initial concentration of trityl dimer 510-3 moll-1)

Conclusions

- A new type of electrospray ionization source for the investigation of thermally initiated reactions by mass spectrometry has been developed.

- Its usefulness has been demonstrated showing two examples. Results are in agreements with calculated reaction kinetics.

- The advantages of this method compared to others are the high sensitivity, specificity  and its speed. Within a few seconds, the user get information about educts, products as well as of reaction intermediates.

Acknowledgement
The authors thank the Deutsche Forschungsgemeinschaft (DFG) for financial support. Travel grants by the DFG and by the German Society for Mass Spectrometry (DGMS) are gratefully acknowledged.

[1]	G. J. van Berkel, The Electrolytic Nature of Electrospray,  in Electrospray Ionization Mass Spectrometry – Fundamentals, Instrumentation and Applications, Ed. by R. B. Cole, Wiley, New York, p. 65-105 (1997)

[2]	J. O. Metzger, J. Griep-Raming, Eur. Mass Spectrom. 1999, 5, 157-163

[3]	J. O. Metzger, Herstellung (Erzeugung) von Radikalen, in Methoden der organischen Chemie (Houben-Weyl), Ed. by M. Regitz, G. Giese, Thieme, Stuttgart, vol. E10a, p. 100