The Persistent Memory Effect of Triethylamine in LC-MS Analysis of Phospholipids

Jens Griep-Raming1, Jürgen O. Metzger1, Heike Rütters2, Thomas Möhring2, Jürgen Rullkötter2

Dept. of Chemistry1, Institute of Chemistry and Biology of the Marine Environment (ICBM)2, both Carl von Ossietzky University of Oldenburg, P. O. Box 2503, D-26111 Oldenburg, Germany

After triethylamine had been used in LC-MS a very intense peak at m/z 102 ([M+H]+ of triethylamine) appeared in the positive ion mode. Only a rigorous cleaning procedure and exchange of plastic parts did remove the contamination. It is recommended not to use triethylamine in LC-MS analysis with quadrupole ion traps.

Triethylamine (TEA) as an eluent modifier for high performance liquid chromatography (HPLC) applications is chosen by chromatographers as an ion pair reagent or as a buffer component of intermediate polarity1. There are examples in the literature that prove its capability to support ionization of various compound classes in electrospray ionization mass spectrometry (ESI-MS)2.

Following Karlsson et al.3, we developed an HPLC/ESI-MS and 
-MS/MS method for the separation and identification of phospholipids from purified bacterial and sediment extracts. Chromatographic separation was carried out on a Thermo Separation Products (TSP, San Jose, CA, USA) HPLC instrument coupled to a Finnigan LCQ ion trap mass spectrometer (Thermoquest-Finnigan, San Jose, CA, USA). Phospholipids were separated according to their headgroups on a LiChrospher 100 Diol 5 HPLC column (125x2mm; Merck, Darmstadt, Germany) using a convex gradient. Eluents contained a formic acid/TEA buffer (0.6/0.08 % v/v, respectively). Measurements in the negative ion mode (spray voltage: -4.5 kV, capillary temperature: 200°C) exclusively yielded [M-H]- ions (except for phosphatidyl choline which formed [M-15]- ions). Figure 1 shows an example of the effective separation and the mass spectra resulting from a single run.

Figure 1 a) TIC chromatogram of PL standards: phosphatidic acid (PA), phosphatidyl glycerol (PG), cardiolipin (CL), phosphatidyl ethanolamine (PE), phosphatidyl choline (PC), phosphatidyl serine (PS), phosphatidyl inositol (PI), all Sigma-Aldrich, Germany ; b) quasi-molecular ions of PI (from soybean) showing different fatty acid combinations (fatty acids are denoted as x:y, with x giving the number of carbon atoms and y representing the number of double bonds, c) MS/MS of m/z 833.

Problems arose when the same instrument was used for measurements in the positive ion mode. Such analyses were virtually made impossible by a very intense signal at m/z 102, [M+H]+ of triethylamine. Direct infusion of pure solvents by the syringe pump resulted in a strongly dominant signal of TEA. The intensity was independent of the kind of solvent supplied. As an ion trap just collects a certain amount of ions and cannot exclude a single m/z species, it was nearly entirely filled with only triethylammonium ions. Therefore, detection of other compounds from diluted solutions was strongly hindered. Only at concentrations as high as 100 µg/mL certain compounds could still be detected. Figure 2, as an example, shows the ESI mass spectrum of 2-(9-decenyl)-5-methyl-1,3,4-oxadiazol at 100 g/mL. Even under these conditions the TEA peak is still significant.

Figure 2 ESI mass spectrum of 2-(9-decenyl)-5-methyl-1,3,4-oxadiazol (ca. 100 µg/mL in methanol)

Even several cleaning cycles of the source and the analyzer did not significantly reduce the intensity of the TEA signal. Cleaning procedures recommended by the instrument manufacturer had no effect on the spectra. Only a rigorous cleaning procedure4 (complete disassembly and factory cleaning) and exchange of most plastic parts removed the TEA contamination to a level below the detection limit.

We strongly recommend refraining from the use of triethylamine in ion-trap LC-MS analysis, and considering alternative eluent modifiers, if the instrument is ever to be used for the analysis of low-molecular weight substances in the positive ion mode.

The authors thank the Deutsche Forschungsgemeinschaft for financial support. A travel grant by the DGMS is gratefully acknowledged.

[1]	V. R. Meyer, Praxis der Hochleistungs-Flüssigchromatographie, 7th ed., Otto Salle Verlag GmbH & Co, Frankfurt am Main, Verlag Sauerländer, Aarau, 1992

[2]	D. E. DeJohn, Proceedings of the 45th ASMS Conference on Mass Spectrometry and Allied Topics, Palm Springs, California, June 1-5, 1997, p. 659.

[3]	A. Å. Karlsson, P. Michelsen, Å. Larsen, G. Odham, Rapid Comm. Mass Spectrom. 1996, 10, 775

[4]	H. Rütters, T. Möhring, J. Rullkötter, J. Griep-Raming, J. O. Metzger, Rapid Commun. Mass Spectrom. 2000, 14, 122–123